## Slide 1: 2020년 8월 회의
한림대학교

## Slide 2
 | activity | 분류 | status | 
 | impedance 실험 | impedance | on going | 
 | PCR speed up | labG | on going | 
 | LoD | labG | on going | o
 | PCR chip taping 고정방법 | labG | on going | 
 | small xavier (3D printer 이용) | labG | on going | 
 | LabGC dialog design | labG | on going | 
신규 | ct threshold 결정 | labG | on going | o
 | shelf motor precision | loco | on going | o
 | SD card chip channel space | loco | on going | 
 | SD card chip cooling speed | loco | on going | 
 | loco HW architecture | loco | on going | 
 | loco SW architecture | loco | on going | 
 | loco battery 용량 실험 | loco | on going | 
 | loco emulator | loco | on going | o
신규 | Fastech python control | loco | on going | o
 | motor-less RT PCR | motorless | on going | o
신규 | rtDuxcycler | rtDux | on going | o

## Slide 3: LabG

## Slide 4
내부 온도 실험을 통한 온도 안정화 및 speedup
50um thermocouple (CHAL-002)  재사용 힘듦
150um (CHAL-005) 로 측정시 온도차이 많음
In-outlet 배치 vs drilled
Thermocouple 대량 구매 요

## Slide 5
PCB taping 효과
Tape O
Tape X
O
X
Cycle 변동이 작다
TC의 실험간 변동 문제 조사중

## Slide 6
d4: PWM on time: 현재 ¼full
Iird4: full PWM + iir filter
시작부분을 보면 상당히 빠를 것을 알 수 있음
parameter tuning 이 필요함  차기 버전에 적용
Speedup
Double sided tape only

## Slide 7
Motor에 의한 mis-location은 무시할 만하다.
Backlash등의 전체 요소를 검증하려면 cartridge 구동부만 open형태 필요
Motor precision

## Slide 8: LoD 실험 및 Ct threshold 계산
황지수교수 발표

## Slide 9: Loco

## Slide 10
Chart: 모터가격 (US$)
| Fastech | Trinamic | w/ encoder | enconder
 | 483.5 | 198.5 | 119 | 27.11
Motor 가격 비교

## Slide 11: Loco emulator
Emulator 2 대 완성
LabG  Loco 안
  Fastech Python control 완성
  LabG 기구 무게 줄임
  Case change

## Slide 12: Motor-less RT PCR

## Slide 13
Find thres sum dist(Abi-labG)
Neural net
10-1
10-2
10-3
10-4
Roche Ct’s
True
Xaiver
Training: 20 curves each
Test: 20 curves each
Cam rtPCR
Training: 20 image seq’s each
Test: 20 image seq’s each
Neural Net을 이용한 Ct calculation

## Slide 14: RT DuxCycler

## Slide 15: 연구항목
형광 검출 검증
duxCycler python control
Raspberry Pi computing module control
Peltier 를 motor driver로 구동
Thermostat를 이용한 lid heater
New camera 선정 (IMX179  IMX415)

## Slide 17
5cm 위에서 center 촬영시 corner에서는 이것보다 더 비스듬히 촬영됨
camera 4대 사용하면 회전 없이 가능함.
Center는 비스듬히 4방향에서 들여다 보고 더함
Side는 2방향에서 비스듬히 들여다 보고 더함.

## Slide 18
Pi computing module
duxCycler python control
Filter wheel python control
4 camera image processing
Flask backend
React frontend

## Slide 19
LED pcb
Ex: 3mm
Em: 5mm
20도 간격배치
빛 차단 sheet
Filter wheel
camera
Filter wheel

## Slide 20
Thermostat 적용검토 (105도 on 107도 off)
Heater에서 가장 먼 곳은 몇 도?

## Slide 21
Peltier driver
Speed to temperature conversion
uStepping과 filter 크기 관계 조사